## Slide 1
Gentle equilibration of macromolecular crystals to new solutions with minimal handling.
Background
Approach
Doug Juers, Sean Stothers,
Jack Quire, Charlie Dibb
Department of Physics & Program in BBMB
Whitman College, Walla Walla WA
Macromolecular crystals are easily damaged during transfer to new solutions due to handling and/or osmotic effects. Here we explore two methods to gently equilibrate crystals to new solutions with minimal handling.
Liquid Based
Aerosol Based
Liquid Based
Aerosol Based
Results & Discussion
Draw vape-pen generated aerosol from syringe
Eject aerosol into cryovial with loaded sample and incubate.
Top: 40 µL pot (pcr tube cap) with inflow and outflow tubes. Bottom: view from below showing alpha lactalbumin crystal on right.
Top: system showing low-cost open-source syringe pumping system (~$400) 1 with enhancements for crystal equilibration. Bottom: auto gradient view of control software
Alpha lactalbumin
0 -> 25 % glycerol
Direct (left) vs 50 min linear gradient (right)
Thermolysin
~2 M AmSO4 -> 0 M Am SO4
Direct (left) vs 15 min linear gradient (right)
Lysozyme
8 % NaCl -> 3% NaCl
Direct (left 2 xtals) vs 40 min linear gradient (right)
References/Acknowledgments
Sina Booeshaghi, Eduardo da Veiga Beltrame, Dylan Bannon, Jase Gehring and Lior Pachter, Principles of open source bioinstrumentation applied to the poseidon syringe pump system, Scientific Reports 9, Article number: 12385 (2019)
This work was supported in part by a grant from the National Institutes of Health (GM0900248).
Thermolysin
~2 M AmSO4 -> 0 M Am SO4
Direct (left) vs ~ 10 min incubation in vial (right)
Batch/Closed System Method
Process/Open System Method
Cryoprotectant is directed towards a loop mounted crystals sitting in a cryovial with a nebulizer plate.
Alpha lactalbumin
0 -> 25 % glycerol
Direct (left) vs ~5 min incubation in vial
Summary
The liquid based approach has proven to be robust and reliable.
The aerosol based approach has been shown to work in principle. The many process variables make it less reliable. However it has potential to be a more rapid method, and further development is underway.
0.5 mm
0.5 mm
0.4 mm
Computer (Raspberry Pi)
Microcontroller
(Arduino)
Microscope
(underneath)
Stepper motor based syringe pumps
3D printed parts
PCR rube cap
(inverted)
crystal
inflow tube
glass coverslip
microscope (underneath)
outflow
inflow
Examples of crystals equilibrated using the described system (right) vs directly transferred to new solution (left).
cryovial
crystal cap
loop mounted crystal (inside vial)
vape-pen
nebulizer plate & new solution (inside)
humid flow
vacuum flow
cryovial end
Loop mounted crystal is inserted into the cryovial end, and incubated in the aerosol under humid flow.